# Part 3 – Technical Requirements

## Supply, Delivery and Installation, of a Point-to-Point Blood Transportation System for Indira Gandhi Memorial Hospital (IGMH) and Dharumavantha Hospital (DH)

## Section 1 – Introduction

To establish an automated specimen transportation network connecting designated clinical locations within IGMH to the Main Laboratory at Dharumavantha Hospital to improve laboratory turnaround time, reduce manual handling, improve specimen traceability and enhance patient care. The need for improved specimen logistics and laboratory turnaround is consistent with the hospital justification and technical evaluation documents.

Bidders shall carefully review this bidding document and submit bids in strict accordance with the instructions below. Failure to submit the required documents, forms, technical evidence, securities, or declarations may result in the bid being considered non-responsive.

- Bidders shall be legally registered entities eligible to supply, install, commission and maintain the proposed system.

- A site visit may be arranged by IGMH. Bidders remain responsible for verifying site conditions, routing constraints, installation requirements and integration needs.

- All documents shall be signed by an authorized representative of the bidder. Any correction, alteration or overwrite shall be initialed. Bidders shall ensure that the submitted technical evidence clearly demonstrates compliance with every requirement in this document.

# Section 2– Scope of Work

The Contractor shall:

### Supply

- Eight (8) Sending Stations

- Four (4) Receiving Stations

- Central Monitoring System

- Transport Network Infrastructure

- Controllers and Communication Equipment

- Software Licenses

- All cabling, tubing, brackets, supports and accessories

### Services

- Site survey

- Engineering design

- Installation

- Testing and commissioning

- Training

- Warranty support

- Preventive maintenance

# Section 3– Technical Specifications

## General System Requirements

The system shall:

- Be a dedicated laboratory specimen transportation system.

- Support direct point-to-point transport.

- One-touch delivery system

- Operate continuously 24/7.

- Support future expansion.

- Include centralized monitoring.

## Performance Requirements

| Requirement | Minimum |
| --- | --- |
| Transport Range | ≥ 1,800 meters |
| Throughput | ≥ 1,000 samples/hour |
| System Availability | ≥ 99% |
| Operation | Fully Automatic |
| Monitoring | Centralized Real-Time |

## Proposed Areas for Installation

Sending Stations

- Main Laboratory (Central Hub/Phlebotomy/OPD Nurses Station) - IGMH

- Emergency Department (Main ER and Pediatric ER)

- Inpatient Ward Area – 1st Floor (Main Theatre Complex, ICU, NIVU, Medical Ward, PVT 2, NCC, Pvt 1, Isolation Ward)

- Inpatient Ward Area – 2nd Floor (Paediatric Ward, Medical Ward Annex, ENT, Ophathalmology, PVR 3, PVT 4)

- Multi-Specialty Surgical Wing

- Hiyalawing Gynae ER

- Hiyalawing Ward A&B, C&D

- Hiyalawing NICU, GOT, LR & LIU

Receiving Station

- Main Laboratory – DH Floor 9 – 4 Receiving stations

## Sending Stations

Quantity: 8

Functions:

- Automated dispatch.

- One touch dispatch.

- Sample verification.

- User interface.

- Status indication.

- Tracking capability.

## Receiving Stations

Quantity: 4

Functions:

- Controlled specimen arrival.

- Safe specimen receiving.

- Arrival notification.

- Laboratory workflow integration.

# Section 4 – Technical Uniqueness Requirements

The offered system shall demonstrate technical innovation supporting laboratory specimen transport.

## Capsule-Free Transport

Preference shall be given to systems capable of direct transport of primary blood collection tubes without dedicated transport capsules.

## Automated Sample Verification

The system should verify dimensional compatibility before dispatch.

## Continuous Transport Operation

The system shall support uninterrupted transport flow and efficient specimen movement.

## Specimen Protection

The system shall incorporate controlled arrival technology and specimen protection mechanisms.

# Section 5– Delivery and Implementation Schedule

| Phase | Duration |
| --- | --- |
| Contract Award | Week 0 |
| Site Survey | 2 Weeks |
| Final Design Approval | 1 Week |
| Manufacturing | 4–8 Weeks |
| Shipping & Clearance | 4–6 Weeks |
| Installation | 4 Weeks |
| Software Configuration | 1 Week |
| Testing & Commissioning | 1 Week |
| Training & Handover | 1 Week |

Maximum Project Duration: 30 Weeks

# Section 6– Training Requirements

The contractor shall provide:

### User Training

- Laboratory staff

- Nurses

- Clinical personnel

### Technical Training

- Biomedical staff

- Maintenance staff

- IT staff

# Section 7– Warranty and Maintenance

Minimum:

- 24 months comprehensive warranty.

- Preventive maintenance.

- Corrective maintenance.

- Software updates.

- Spare parts support.

Spare part availability shall be maintained for at least ten years.

# Section 8 – Technical Evaluation

| Criteria | Marks |
| --- | --- |
| Compliance with Technical Specifications | Mandatory |
| System Performance | Mandatory |
| Warranty & Maintenance | Mandatory |
| Training & Technology Transfer | Mandatory |
| Local Support Capability | Mandatory |

Section 9

The following shall be completed and submitted by bidders as part of the bid. Where a document is not applicable, the bidder shall clearly state “Not Applicable” and provide justification.

- Bid Submission Form

- Price Quotation

- Completed Compliance Matrix

- Bid Security Document

- Last 3 years Audited Financial Statements

- Manufacturer Authorization Letter

- Experience Letter / Details of the implement bases

- Warranty Commitment letter

- Service Support Declaration with details of trained engineers

- Proposed Implementation Schedule

- Staff Deployment Plan

Bid Submission Form

To: Indira Gandhi Memorial Hospital. We, the undersigned, offer to supply, deliver, install, test, commission, train, warrant and maintain the Point-to-Point Blood Transportation System in accordance with the bidding document. We confirm that our bid is valid for 120 days and that all information submitted is true and complete.

| Bidder Name | ______________________________ |
| --- | --- |
| Authorized Representative | ______________________________ |
| Designation | ______________________________ |
| Signature and Seal | ______________________________ |
| Date | ______________________________ |

Form 2 – Compliance Matrix

| No. | Requirement | Comply / Not Comply | Evidence / Page Reference | Remarks |
| --- | --- | --- | --- | --- |
| 1 | System supports point-to-point specimen transport. |  |  |  |
| 2 | Minimum transport range of 1,800 meters. |  |  |  |
| 3 | Eight sending stations and four receiving stations supplied. |  |  |  |
| 4 | Central monitoring and real-time tracking included. |  |  |  |
| 5 | Once-Touch Delivery |  |  |  |
| 6 | Capsule free Transport |  |  |  |
| 7 | Automated Sample Verification |  |  |  |
| 8 | Continuous Transport Operation |  |  |  |
| 9 | Specimen Protection |  |  |  |
| 10 | Service Support |  |  |  |
| 11 | Proposed Implementation Plan |  |  |  |
| 12 | Staff Training / |  |  |  |
| 13 | Minimum 24-month comprehensive warranty provided. |  |  |  |

Conclusion

The procurement of a Point-to-Point Blood Transportation System for Indira Gandhi Memorial Hospital and Dharumavantha Hospital is an important initiative to improve the efficiency, safety, traceability and reliability of laboratory specimen transportation across key clinical service areas. The proposed system is expected to reduce manual handling, minimize delays in specimen movement, improve laboratory turnaround time, and support timely clinical decision-making for patient care.

This bidding document has been prepared to provide clear instructions to bidders, define the required scope of supply and services, establish minimum technical and performance requirements, and set out the evaluation, award, warranty, training, implementation and submission requirements. The document is intended to ensure that all bids are assessed in a transparent, consistent and technically sound manner, in line with the requirements of the Procuring Entity and the applicable procurement procedures.

Accordingly, bidders are required to submit complete, compliant and evidence-based proposals demonstrating their capacity to supply, install, test, commission, train, warrant and maintain the proposed system within the specified timeframe. The successful implementation of this project will strengthen hospital logistics, enhance laboratory workflow, and contribute to improved quality and responsiveness of healthcare services at IGMH and Dharumavantha Hospital.